README for the software provided along the paper:

M.Vulovic, B. Rieger, L.J. van Vliet, A.J. Koster, R.B.G. Ravelli,

A Toolkit for the Characterisation of CCD Cameras for Transmission Electron Microscopy, Acta Cryst. (2010). D66, 97-109  [ doi:10.1107/S0907444909031205 ].

The algorithms to characterise CCD detectors have been implemented in DipImage, a publicly available software toolbox ( www.diplib.org ) for Matlab (The MathWorks, Inc.)

The following images are required:

1) Dark reference images (with the column valves closed, i.e. there was no beam)

2) White reference images (with uniform illumination of the detector)

3) The normalized images of a beam stop or diffraction aperture (MTF edge method)

The images should be saved with the following name convention:

1) Bin%d_Exp%05.2f_dark%02d for dark reference images

2) Bin%d_Exp%05.2f_gain%02d  for white reference images

where %d and %05.2f stand for the syntax of printf in C or Matlab, .e.g.

Bin2_Exp01.00_white01 for binning 2, exposure time of 1 second and white image number 1.

Have a look at the file example1.m which shows the processing of the provided images.

Changes in Version 2 of the TEM CCD software:

The function ex_slice was changed to be compatible with Dipimage 2.3. The first version of  TEM CCD software works with all Dipimage versions before 2.3.

|  | FUNCTION (m.file) | Description |
| --- | --- | --- |
| 1 | remout | Remove outliers and writes output images in the same folder |
| 2 | lsq | Makes a linear least square fit of the intensity against integration (exposure) times of each pixel in a stack of images; computes offset and slope of the fit |
| 3 | average | Calculates an average background or white reference image for a certain exposure time and binning |
| 4 | averagesame | Averages the images within the stack with the same exposure time; writes out the average image |
| 5 | corr_raw | Corrects a raw image for the bias and gain |
| 6 | ef_gain | Calculates average variance and average intensity within stacks of uniformly illuminated images with different exposure times; plots the effective detector gain |
| 7 | r_square | Calculates the coefficient of determination to check how well the estimated regression line fits the data |
| 8 | readout_noise_pix | Calculates the readout noise image from the stack of dark images acquired with the lowest possible integration time and writes it out |
| 9 | mtf_noise | MTF via noise method |
| 10 | mtf_edge | Calculates MTF via straight edge (beam stop) |
| 11 | mtf_edgeAperture | Calculates MTF via diffraction aperture |
| 12 | dqe | Calculates Detector Quantum Efficiency |

24.08.2011.